Upgrade your Master-9:

- Turn off the Master-9 unit and remove its top cover.

- On the circuit board, switch the DOWNLOAD switch to position “1” (GND) see the following picture:

- Turn on the Master-9 and connect it to the PC via the USB cable.

- For the following steps, use the files that you saved in the Master-9 folder.

If the serial number of your Master-9 is lower than 5325 (lower than hardware_6.9 and software_7.6), you MUST first install the hf6_8.flash and the sf6_8.flash.

If the serial number of your Master-9 is higher than 5325 the hf6_8.flash and the sf6_8.flash are already installed in your unit and you can skip to step # 5.

hf6_8.flash installation:
Run the Master9_updater.exe.
Browse to the "hf6_8.flash " file.
Press: START.
The installation takes about 10 minutes.

sf6_8.flash installation:
Run the Master9_updater.exe.
Browse to the "sf6_8.flash " file.
Press: START.
The installation takes about 2 minutes.

5. Hardware and Software upgrades:

Hardware upgrade:
Run the Master9_updater.exe.
Browse to the "hardware_app_ver10_5.flash" file.
Press: START.
The upgrading takes about 10 minutes.

Software upgrade:
Run the Master9_updater.exe.
Browse to the "software_app_ver11_16.flash" file.
Press: START.
The upgrading takes about 2 minutes.

- Turn off the Master-9 and switch back the DOWNLOAD switch to position “2”.

- Turn on the Master-9. On the touch screen of the Master-9 press the VER key and see that the new hardware and software versions are installed.